## Slide 1: Radiofrequency Pulses in MRI
Mahesh Keerthivasan
BME 639
University of Arizona
Lecture 1

## Slide 2: RF Excitation
Adapted from Larry Wald’s slides, fMRI MR Physics
Main
Field
Bo
x
y
z
North
M0
Applied B1 RF Field
B1(t)
x
y
z
M0
The RF pulse rotates Mo the about applied field
B1

## Slide 3: The B1 Field
B1 field: a time varying radio frequency field at Larmor frequency and perpendicular to B0
B0
Alternating current at Larmor frequency
Birdcage RF Coil
B1 field
http://mriquestions.com/

## Slide 4: The B1 Field
B1 field generation:
Frequency synthesizer
(generate carrier wave at Larmor frequency)
Pulse Modulator
(use short RF pulses to modulate the carrier wave)
Power Amplifier
(control peak B1 field amplitude and flip angle)
Transmit Coil
(quadrature or circularly polarized to generate field)
http://mriquestions.com/
Sinusoidal carrier wave at
Larmor frequency >~48MHz
Carrier frequency modulated by a low frequency sinc RF ~kHz
RF Envelope B1(t)

## Slide 5: The B1 Field
A circularly polarized B1 field at carrier frequency is represented in the laboratory frame:
Transforming to the rotating frame and using as the angular rotating frequency :
=
The rotating frame demodulates the RF oscillations
into the envelope
Handbook of MRI Pulse Sequences, Bernstein
http://mriquestions.com
Laboratory frame
Rotating frame

## Slide 6: Motivation
We would like to understand:
Properties of the RF pulse
Effect of RF on magnetization
Design efficient RF pulses
time

## Slide 7: Flip Angle
Magnetization is tipped out of equilibrium due to applied B1 field
x
y
z
M0
Bo
B1
Flip angle when RF frequency matches Larmor frequency:
Flip angle is the area under envelope of RF field

## Slide 8: Flip Angle
Flip angle = 30o
Flip angle = 45o
Flip angle = 90o

## Slide 9: Slice Selective Excitation
Every pulse sequence you have seen so far has a RF pulse along with a gradient waveform
BME639, Dr. Nan-kuei Chen

## Slide 10: Slice Selective Excitation
By applying a slice selection gradient along with the RF pulse we can localize excitation to a specific slice or region.
Handbook of MRI Pulse Sequences, Bernstein
Slides from Dr. Nan-kuei Chen
Notes: - The RF pulse and the slice select gradient strength determine the location and thickness of the slice

## Slide 11: Slice Selective Excitation
Slice select gradient varies Larmor frequency along slice direction
https://www.radiologycafe.com/radiology-trainees/frcr-physics-notes/
Notes: - The RF pulse and the slice select gradient strength determine the location and thickness of the slice

## Slide 12: Slice Selective Excitation
Bo
Slice select gradient varies Larmor frequency along slice direction
Slides from Dr. Nan-kuei Chen
Notes: - The RF pulse and the slice select gradient strength determine the location and thickness of the slice

## Slide 13: Slice Selective Excitation
This allows RF pulse to selectively excite spins by choosing center frequency
Slides from Dr. Nan-kuei Chen
Bo
RF
Notes: - The RF pulse and the slice select gradient strength determine the location and thickness of the slice

## Slide 14: Slice Selective Excitation
Bo
RF
This allows RF pulse to selectively excite spins by choosing center frequency
Slides from Dr. Nan-kuei Chen
Notes: - The RF pulse and the slice select gradient strength determine the location and thickness of the slice

## Slide 15: Slice Selective Excitation
https://www.radiologycafe.com/radiology-trainees/frcr-physics-notes/
This allows RF pulse to selectively excite spins by choosing center frequency
Notes: - The RF pulse and the slice select gradient strength determine the location and thickness of the slice

## Slide 16: Slice Selective Excitation
Gz
frequency
Thickness of the excited slice

## Slide 17: Slice Selective Excitation
Gz2
Gz1
frequency
Thickness of the excited slice
In practice, the RF bandwidth is fixed and gradient strength varied

## Slide 18: Slice Selective Excitation
Adapted from MRI from Picture to Proton. Cambridge: Cambridge University Press
Changing the slice location by changing the RF center frequency

## Slide 19: Quick Summary
High frequency carrier wave modulated by low frequency RF pulse to generate B1 field
Flip angle : controls tip angle of magnetization
RF pulse and slice select gradient enable localization of excitation
Slice thickness controlled by RF Bandwidth and gradient amplitude

## Slide 20: Effect of RF Pulse on Magnetization
A selective RF Pulse excites spins in the slice of interest
Ideally we would like the magnetization from all spins in the selected slice to be tipped at the same flip angle
Flip angle
Ideal slice profile
Slice position
Slice profile  flip angle distribution across slice
Notes: - Want uniform flip angle within passband of RF

## Slide 21: Effect of RF Pulse on Magnetization
Can we measure the flip angle distribution across the slice ?
How does the RF pulse affect magnetization in the slice of interest ?
Notes: - Want uniform flip angle within passband of RF

## Slide 22: Effect of RF Pulse on Magnetization
Bloch Equations: To model effect of B1 field on magnetization
Assuming no T1, T2 relaxation and RF pulse applied in direction :
Notes: - Want uniform flip angle within passband of RF

## Slide 23: Effect of RF Pulse on Magnetization
Can be solved assuming RF frequency is the Larmor frequency () :
    Where flip angle
    Transverse magnetization
Handbook of MRI Pulse Sequences, Bernstein
Notes: - Want uniform flip angle within passband of RF

## Slide 24: Effect of RF Pulse on Magnetization
Since spins are on resonance only at center of the slice,
The transverse magnetization can be modeled by
Results in the following solution
Inverse Fourier Transform of
Notes: - Want uniform flip angle within passband of RF

## Slide 25: Effect of RF Pulse on Magnetization
If the flip angle is small and
    It can be shown that,
    and for small angle the flip angle distribution is
Inverse Fourier Transform of
Notes: - Want uniform flip angle within passband of RF

## Slide 26: RF Pulse Slice Profile
Small tip angle approximation:
Slice profile is approximately equal to inverse Fourier transform of RF pulse
Fourier Transform
Slice profile  flip angle distribution across slice
https://www.chem.purdue.edu/courses/chm621/text/ft/basiset/rect/rect.htm

## Slide 27: RF Pulse Slice Profile
Small tip angle approximation:
Slice profile is approximately equal to inverse Fourier transform of RF pulse
Valid for , breaks down for
Slice profile  flip angle distribution across slice
Handbook of MRI Pulse Sequences, Bernstein

## Slide 28: RF Pulse Slice Profile
Small tip angle approximation:
Slice profile is approximately equal to inverse Fourier transform of RF pulse
Valid for , breaks down for
Slice profile  flip angle distribution across slice
Handbook of MRI Pulse Sequences, Bernstein
Calculate the slice profile by solving the Bloch equations !